## USB_DataLogger (old) (hidden)
JSON Schema
Database structure
| Tag/Key | Decription | Unit | SQL Type | Size
| ID | Device ID assigned to USB data logger |  | TINYINT | 1 byte
| Time | Timestamp of when Data log is received by the websever |  | DATETIME | 5 bytes + fractional seconds storage
| Penning_Pressure_Start | Penning pr (Pa) at time = 0 | Pascals | FLOAT | 4 bytes
| MSK_Pressure_Start | MSK (Pa) at time = 0 | Pascals | FLOAT | 4 bytes
| Penning_Pressure_Start | Penning pr (Pa) at time = max (i.e. end of log) | Pascals | FLOAT | 4 bytes
| MSK_Pressure_End | MSK (Pa) at time = max (i.e. end of log) | Pascals | FLOAT | 4 bytes
| ICP_Power_Start_Time | Process time: calculate the time when column "ICP power (W)"> 0 | ms | INT | 4 bytes
| ICP_Power_Average | The average value of "ICP power (W)" when "ICP power (W)"> 0 | Watts | FLOAT | 4 bytes
| Cleaning | If the recipe contains in the file name "StandardCleaning", mark it as "cleaning" |  | TINYINT(1) | 1 byte
Each Data Log object should contain the following information
Column Name | Tag/Key | Description | Unit | Type | Size
Penning pr. (Pa) | Penning_Pressure_Start | Penning pr (Pa) at time = 0 | Pascals | FLOAT | 4 BYTES
MSK (Pa) | MSK_Pressure_Start | MSK (Pa) at time = 0 | Pascals | FLOAT | 4 BYTES
Penning pr. (Pa) | Penning_Pressure_Start | Penning pr (Pa) at time = max (i.e. end of log) | Pascals | FLOAT | 4 BYTES
MSK (Pa) | MSK_Pressure_End | MSK (Pa) at time = max (i.e. end of log) | Pascals | FLOAT | 4 BYTES
ICP power (W) | ICP_Power_Start_Time | Process time: calculate the time when column "ICP power (W)"> 0 | ms | INT | 4 BYTES
ICP power (W) | ICP_Power_Average | The average value of "ICP power (W)" when "ICP power (W)"> 0 | Watts | FLOAT | 4 BYTES
File name | Cleaning | If the recipe contains in the file name "StandardCleaning", mark it as "cleaning" |  | BOOL | 1 BIT

## 0x10-Transmit
Digi Xbee Zigbee 3.0 API
0x10 - Transmit Request (pg. 311)
Frame structure:
Index | 0 | 1 | 3 | 4 | 5 | 13 | 15 | 16 | 17 | 17+L
Length (Bytes) | 1 | 2 | 1 | 1 | 8 | 2 | 1 | 1 | L | 1
Field | <Start Deliminator> | <length> | <Frame type> | <Frame ID> | <64-bit dest. address> | <16-bit dest. address> | <broadcast radius> | <options> | <YOUR DATA STRUCTURE> | <Checksum>
Fields:
Start Delimitator | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | Indicates the start of the frame - Always equal to "0x7E" | [6 empty] | 7E
Length | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 2 | The number of bytes between the length and checksum. For a data structure of L bytes, the length value is L + 14 | [6 empty] | 00 27 (For data structure of 25 bytes, the length is: 25 + 14 = 39 = 0x27)
Frame Type | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | Type of Digi API Frame - Transmit Request - 0x10 | [6 empty] | 10
Frame ID | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | Identifies the data frame for the host to correlate with a subsequent response frame. If set to 0, the device will not emit a response frame. | [6 empty] | 01
64 Bit Destination Address | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 8 | The MAC address of the destination Device. When sending to the network coordinator, use all zeros (0x0000000000000000) | [6 empty] | 00 00 00 00 00 00 00 00
16 Bit Destination Address | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 2 | The 16-bit network address of destination device. Set to 0xFFFE to use 64-bit addressing instead. | [6 empty] | FF FE
Broadcast Radius | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | Max number of hops for broadcast transmission (not required). Set to zero to use the NH parameter (set in firmware config) | [6 empty] | 00
Options | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | If equal to zero, use of the firmware parameter "TO" is used. Bit Map as follows: | [6 empty] | 00
|  | Bit | Meaning |  | Description
|  | 0 | Disable ACK |  | Description
|  | 1 | Reserved |  | Description
|  | 2 | Indirect Transmission |  | Description
|  | 3 | Multicast |  | Description
|  | 4 | Secure Session Encryption |  | Description
|  | 5 | Enable APS Encryption |  | Description
|  | 6 | Use extended timeouy |  | Description
YOUR DATA STRUCTURE | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| L | The Zigbee dataframe payload (of length L bytes) | [6 empty] | 37 0B 19 93 47 02 79 00 38 99 72 D8 3C AB 8E 8F 00 02 4C D7 44 38 C0 00 01
Checksum | Length (Bytes) | Description | [6 empty] | Example (Bytes - Hexadecimal)
| 1 | The frame checksum. This value will need to be calculated for each frame. It is calculated as 0xFF minus the 8-bit sum of bytes from the <Frame type> until this field (i.e. the whole frame excluding the <Start Delimiter>, the <Length> and the <Checksum>) | [6 empty] | C0 For the total frame formed from all above example values
Full Zigbee Frame Example |  | 7E 00 27 10 01 00 00 00 00 00 00 00 00 FF FE 00 00 37 0B 19 93 47 02 79 00 38 99 72 D8 3C AB 8E 8F 00 02 4C D7 44 38 C0 00 01 C0

## 0x11-Explicit-Transmit
Digi Xbee Zigbee 3.0 API
0x11 - Explicit Addressing Command Request
This frame type is used to send payload data as an RF packet to a specific destination 0x10, but with additional fields available for user-defined endpoints, cluster ID, and profile ID.
Digi.xbee Library API
Class "XBeeDevice", importable from digi.xbee.devices includes the method:
remote_addr = XBee64BitAddress.from_hex_string("0013A200XXXXXXXX")# Replace with the 64-bit address of the power outlet remote_device = RemoteXBeeDevice(device, remote_addr) # Zigbee On command for On/Off cluster (this is a simplified payload) SOURCE_ENDPOINT = 0x01 # This can be almost any arbitrary endpoint for the source DEST_ENDPOINT = 0x01 # As mentioned, this is a common default, but it might vary CLUSTER_ID = 0x0006 #ON/OFF Zigbee cluster PROFILE_ID = 0x0104 #Home Automation Zigbee Profile ID OFF_COMMAND = bytes([0x00]) # This represents the "Off" command in ZCL for On/Off cluster ON_COMMAND = bytes([0x01]) # This represents the "On" command in ZCL for On/Off cluster #Function to turn off power point device.send_data_explicit(remote_device, SOURCE_ENDPOINT, DEST_ENDPOINT, CLUSTER_ID, PROFILE_ID, OFF_COMMAND) #Function to turn on power point device.send_data_explicit(remote_device, SOURCE_ENDPOINT, DEST_ENDPOINT, CLUSTER_ID, PROFILE_ID, ON_COMMAND) # Note: This assumes the device you're sending to is using endpoint 0x01 for its On/Off cluster. # For a more robust solution, you would perform endpoint and cluster discovery as mentioned.
API Reference for "digi-xbee" python library
https://xbplib.readthedocs.io/en/latest/user_doc/communicating_with_xbee_devices.html#send-explicit-data-to-one-device
xbee.send_expl_data(remote_device, SOURCE_ENDPOINT, DEST_ENDPOINT, CLUSTER_ID,PROFILE_ID,COMMAND)

## 0x90-Recieve
Digi Xbee Zigbee 3.0 API
0x90 - Receive Packet (pg. 345)
Frame received when data has been sent by a device using the 0x10 (or 0x11) Digi API command
Frame structure:
Index | 0 | 1 | 3 | 4 | 12 | 14 | 15 | 15+L
Length (Bytes) | 1 | 2 | 1 | 8 | 2 | 1 | L | 1
Field | <Start Deliminator> | <length> | <Frame type> | <64-bit source address> | <16-bit source address> | <options> | <RECEIVED DATA STRUCTURE> | <Checksum>
Fields:
Start Deliminator | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 1 | Indicates the start of the frame - Always equal to "0x7E" |  |  |  |  | 7E
Length | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 2 | The number of bytes between the length and checksum. For a data structure of L bytes, the length value is L + 13 |  |  |  |  | 00 27 (For data structure of 25 bytes, the length is: 25 + 14 = 39 = 0x27)
Frame Type | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 1 | Type of Digi API Frame - Receive Packet - 0x90 |  |  |  |  | 90
64 Bit Source Address | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 8 | The sending device's MAC address. |  |  |  |  | 00 00 00 00 00 00 00 00
16 Bit Destination Address | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 2 | The sending device's 16-bit network address. |  |  |  |  | FF FE
Options | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 1 | Bit Map as follows: |  |  |  |  | 00
|  | Bit | Meaning
|  | 0 | Package was Acknowledged
|  | 1 | Package was sent as a broadcast
|  | 2 | Reserved
|  | 3 | Reserved
|  | 4 | Packet was sent across a secure session
|  | 5 | Package was encrypted with Zigbee APS security
|  | 6,7 | See manual
YOUR DATA STRUCTURE | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| L | The received data structure of length L |  |  |  |  | 37 0B 19 93 47 02 79 00 38 99 72 D8 3C AB 8E 8F 00 02 4C D7 44 38 C0 00 01
Checksum | Length (Bytes) | Description |  |  |  |  | Example (Bytes - Hexadecimal)
| 1 | The frame checksum. This value will need to be calculated for each frame. It is calculated as 0xFF minus the 8-bit sum of bytes from the <Frame type> until this field (i.e. the whole frame excluding the <Start Delimiter>, the <Length> and the <Checksum>) |  |  |  |  | C0 For the total frame formed from all above example values
Full Zigbee Frame Example |  | 7E 00 27 10 01 00 00 00 00 00 00 00 00 FF FE 00 00 37 0B 19 93 47 02 79 00 38 99 72 D8 3C AB 8E 8F 00 02 4C D7 44 38 C0 00 01 C0

## Interlock
Interlock Device
Summary
|  | 1. | When someone tries to log into the interlock device (either with user credentials of with their RFID), the interlock device will send the "Verification" Zigbee data frame structure to the gateway device using the Digi API "0x10 Transmit request" command.
|  | 2. | Gateway will receive data frame with Digi API function "0x90 - Receive Packet". The gateway device will then check the credentials (or RFID) against the local acess database to check if the user has access to corresponding interlock module.
|  | 3. | The gateway will then send the "Validate" Zigbee data frame structure back to the interlock device using the Digi API "0x10 Transmit request" command (with the interlock device recieving a data frame with Digi API function "0x90 - Receive Packet". If the verification was successful, the gateway must send the "Switch" Zigbee dataframe using the Digi API "0x11 Explicit Addressing Command Request" to turn on the corresponding power switch.
|  | 4. | The gateway will send a log of the verification attempt to the Websever using the JSON API.
|  | 5. | Before the time expires, the gateway still needs to check for messages from the interlock device as it may receive a "Verify" Zigbee dataframe from the interlock device, either with command "Logout" (0x00) or "Extend" (0x03). In the case of receiving command "Extend" (0x03), add the extended time to the timer for the corresponding power switch.
|  | 6. | At the end of the timer or upon reciept of a "Verification" frame from the interlock device with command "Logout" (0x00) , the gateway must send the "Switch" Zigbee dataframe using the Digi API "0x11 Explicit Addressing Command Request" to turn off the corresponding power switch.
|  | 7. | After turning the power switch off, the gateway must send a log of the turn off to the Webserver via the JSON API
"Verify" Zigbee Data Structure (From interlock to gateway)
Sent from an interlock device to the gateway. If the frame is of type 0x01 or 0x02, a "Validation" frame is expected in return from the gateway.
|  | First Byte Index | Length (bytes) | Type | Name | Description
|  | 0 | 1 | bitmap | Command | See Bitmap description below
|  | 1 | variable | Command Specific payload structure below
Command Value - '0x00' - "Logout"
No Payload
Command Value - '0x01' - "Credential Login"
|  |  | Length (bytes) | Type | Name | Description
|  | 1 | 4 | uint32 | User ID | The user ID entered by the user attempting to login
|  | 5 | 4 | char | Pin | 4 digit user pin (4 characters)
|  | 9 | 1 | uint8 | Time | The time (in minutes up to 255), requested to have the interlock device turned on for.
Command Value - '0x02' - "RFID Login"
|  |  | Length (bytes) | Type | Name | Description
|  | 1 | 4 | byte | RFID | The RFID of the user attempting to login
|  | 5 | 1 | uint8 | Time | The time (in minutes up to 255), requested to have the interlock device turned on for.
Command Value - '0x03' - "Extend"
|  |  | Length (bytes) | Type | Name | Description
|  | 1 | 1 | uint8 | Time | The time (in minutes up to 255), requested to have the interlock device turned on for.
"Validate" Zigbee Data Structure (From gateway to interlock upon reciept of Verification")
Sent from the gateway to the interlock device after receiving a "Verify" frame with command "Credential Login" (0x01) or "RFID Login" (0x02)
|  | First Byte Index | Length | Type | Name | Description
|  | 0 | 1 | psuedo boolean | Outcome | The outcome of the login attempt. If unsuccessful, Outcome is equal to "0x00" If successful, outcome is equal to "0x01"
|  | Total Length | 1 | Bytes

## ON-OFF_Cluster (hidden)
On/Off Zigbee Cluster |  |  | (See pg 3-61 of Zigbee Cluster Library Specification V8 or pg. 307 of NPX Zigbee Cluster Library (for ZigBee 3.0)
Cluster ID | 6
Attributes | [6 empty] | Commands
Identifier | Name | Type | Range | Acc | Def | M |  | ID | Description | M/O
0x0000 | OnOff | bool | value | RPS | 0 | M |  | 0x00 | Off | M
0x4000 | GlobalSceneControl | bool | value | RPS | 1 | O |  | 0x01 | On | M
0x4001 | OnTime | unint16 | full-non | RW | 0 | O |  | 0x02 | Toggle | M
0x4002 | OffWaitTime | uint16 | full-non | RW | 0 | O |  | 0x40 | Off with effect | O
0x4003 | StartUpOnOff | enum8 | desc | RW | MS | O |  | 0x41 | On with recall global scene | O
[8 empty] | 0x42 | On with timed off | O

## RS232
RS232 Data Logger
Summary
|  | 1. | After completing a full log, the RS232 data logger will begin transmission of the whole log. The beggining of a log will begin with the transmission of Zigbee Data Structure 1 (header frame), followed by multiple Zigbee Data Structure 2 (body frames) until the whole file has been transmitted. |  |  |  | Example of Zigbee Data Structure 1
|  | 2. | Gateway will receive data packet with Digi API frame of function "0x90 - Receive Packet". By checking the length of the data received, the gateway can determine which Zigibee data structure was sent, and thus whether to create a new log, or append the contents to an existing log. |  |  |  | Original Data: | Density = 1.23 (3F 9D 70 A4) Tooling = 150.8 (43 16 CC CD) Z-Ratio = 42.9 (42 2B 99 9A)
|  | 3. | Gateway will convert the received zigbee data, and write it to a .csv or .txt file. (TBC) The gateway will transfer the file to the webserver using either JSON or a file transfer protocol [will be HTTPS request]. The gateway must also indicate to the webserver the time the header frame was sent and the MAC address (or Device ID) of the device which sent the data. |  |  |  | Bytes Sent: | 3F 9D 70 A4 43 16 CC CD 42 2B 99 9A
|  | 4. | The server stores the log data.
Example of Zigbee Data Structure 2
| Zigbee Data Structure 1 (Header frame) | [5 empty] | Time (s) | Value | Recorded Data (Float) | Recorded Data (Bytes)
|  | First Byte Index | Length | Type | Name | Description | 1 | Rate | 4.53 | 40 90 F5 C3
|  | 0 | 4 | float | Density | Density of material |  | Thickness | 1.42 | 3F B5 C2 8F
|  | 4 | 4 | float | Tooling | The correction needed between the thickness of the layer deposited on the crystal monitor and the thickness of the film deposited on the substrate | 2 | Rate | 4.53 | 40 90 F5 C3
|  | 8 | 4 | float | Z_Ratio | Z-ratio of material |  | Thickness | 1.42 | 3F B5 C2 8F
|  | Total Length | 12 | Bytes |  |  | 3 | Rate | 4.53 | 40 90 F5 C3
[8 empty] | Thickness | 1.42 | 3F B5 C2 8F
| Zigbee Data Structure 2 (Body frame)* | [5 empty] | 4 | Rate | 4.53 | 40 90 F5 C3
|  | First Byte Index | Length | Type | Name | Description |  | Thickness | 1.41 | 3F B4 7A E1
|  | 0 | 4 | float | Rate1 | Rate of deposition on the film Sample 1 second later | 5 | Rate | 4.53 | 40 90 F5 C3
|  | 4 | 4 | float | Thickness1 | Thickness of film Sample 1 second later |  | Thickness | 1.41 | 3F B4 7A E1
|  | 8 | 4 | float | Rate2 | Rate of deposition on the film Sample 1 second later | 6 | Rate | 4.53 | 40 90 F5 C3
|  | 16 | 4 | float | Thickness2 | Thickness of film Sample 1 second later |  | Thickness | 1.4 | 3F B3 33 33
|  | … | … | … | … | … | 7 | Rate | 4.2 | 40 86 66 66
|  | 184 | 4 | float | Rate24 | Rate of deposition on the film Sample 1 second later |  | Thickness | 1.39 | 3F B1 EB 85
|  | 188 | 4 | float | Thickness24 | Thickness of film Sample 1 second later | 8 | Rate | 4.2 | 40 86 66 66
|  | 192 | 4 | float | Rate25 | Rate of deposition on the film Sample 1 second later |  | Thickness | 1.39 | 3F B1 EB 85
|  | 196 | 4 | float | Thickness25 | Thickness of film Sample 1 second later | 9 | Rate | 4.2 | 40 86 66 66
|  | Total Length | 200 | Bytes |  |  |  | Thickness | 1.39 | 3F B1 EB 85
[7 empty] | 10 | Rate | 4.2 | 40 86 66 66
|  | *NOTE | There is currently no limit on the maximum length of the log so the gateway will keep receiving frames with 25 seconds worth of data (one body frame) until the the whole log has been sent by the RS-232 module. There is no time information transferred over zigbe. When the gateway is writing the data to a .csv it should include the time series (e.g. increment every second and save this for each Rate & Tooling pair). |  |  |  |  | Thickness | 1.38 | 3F B0 A3 D7
[7 empty] | 11 | Rate | 4.2 | 40 86 66 66
[8 empty] | Thickness | 1.33 | 3F AA 3D 71
[7 empty] | 21 | Rate | 2 | 40 00 00 00
[8 empty] | Thickness | 1.33 | 3F AA 3D 71
[7 empty] | 22 | Rate | 2 | 40 00 00 00
[8 empty] | Thickness | 1.32 | 3F A8 F5 C3
[7 empty] | 23 | Rate | 2 | 40 00 00 00
[8 empty] | Thickness | 1.32 | 3F A8 F5 C3
[7 empty] | 24 | Rate | 2 | 40 00 00 00
[8 empty] | Thickness | 1.32 | 3F A8 F5 C3
[7 empty] | 25 | Rate | 2 | 40 00 00 00
[8 empty] | Thickness | 1.32 | 3F A8 F5 C3
[7 empty] | Bytes Sent: | 40 90 F5 C33F B5 C2 8F40 90 F5 C33F B5 C2 8F40 90 F5 C33F B5 C2 8F40 90 F5 C33F B4 7A E140 90 F5 C33F B4 7A E140 90 F5 C33F B3 33 3340 86 66 663F B1 EB 8540 86 66 663F B1 EB 8540 86 66 663F B1 EB 8540 86 66 663F B0 A3 D740 86 66 663F AA 3D 7140 00 00 003F AA 3D 7140 00 00 003F A8 F5 C340 00 00 003F A8 F5 C340 00 00 003F A8 F5 C340 00 00 003F A8 F5 C3

## USB
USB Data Logger
Summary
|  | 1. | Transmission initiated by USB logger sending a data packet to the gateway using Digi API function "0x10 - Transmit Request"
|  | 2. | Gateway will receive data packet using Digi API function "0x90 - Receive Packet"
|  | 3. | Gateway will convert the received zigbee data, and place it in a dictionary to send to webserver using JSON API. This dictionary should include the time that the data was received, and the MAC address (or Device ID) of the device which sent the data.
|  | 4. | The server stores the log data.
Zigbee Data Structure
|  | First Byte Index | Length | Type | Name | Description
|  | 0 | 4 | float | Penning_Pressure_Start | Penning pr (Pa) at time = 0
|  | 4 | 4 | float | MSK_Pressure_Start | MSK (Pa) at time = 0
|  | 8 | 4 | float | Penning_Pressure_Start | Penning pr (Pa) at time = max (i.e. end of log)
|  | 12 | 4 | float | MSK_Pressure_End | MSK (Pa) at time = max (i.e. end of log)
|  | 16 | 4 | float | ICP_Power_Start_Time | Process time: calculate the time when column "ICP power (W)"> 0
|  | 20 | 4 | float | ICP_Power_Average | The average value of "ICP power (W)" when "ICP power (W)"> 0
|  | 24 | 1 | byte | Cleaning | If the recipe contains in the file name "StandardCleaning", mark it as "cleaning"
|  | Total Length | 25 | Bytes
JSON Data Structure
|  |  | Key | Value Type | Name | Description
|  |  | dev_id | int | device_id | the reference number to the corresponding device for identification
|  |  | log_data | JSON | log_data | Nested JSON file containing the data to be logged
|  |  | timestamp_str | string | time submitted | timestamp of when the data was logged on web server
|  |  | frame | int | dataframe | The reference number to identify which data frame is being sent to the web server, so that the data can be correctly logged
Inside nested JSON (log_data)
| If device id = 2 & frame = '000' |  | pen_s | float | Penning_Pressure_Start | Penning pr (Pa) at time = 0
|  |  | MSK0 | float | MSK_Pressure_Start | MSK (Pa) at time = 0
|  |  | pen_m | float | Penning_Pressure_Start | Penning pr (Pa) at time = max (i.e. end of log)
|  |  | MSKM | float | MSK_Pressure_End | MSK (Pa) at time = max (i.e. end of log)
|  |  | p_time | float | ICP_Power_Start_Time | Process time: calculate the time when column "ICP power (W)"> 0
|  |  | ICP | float | ICP_Power_Average | The average value of "ICP power (W)" when "ICP power (W)"> 0
|  |  | cleaning | int | cleaning | If the recipe contains in the file name "StandardCleaning", mark it as "cleaning"

## RS232_DataLogger(old) (hidden)
| Variable | Description | Unit | Type | Size
1.11 | Rate | Rate of deposition on the film | A/s | FLOAT | 4 Bytes
0.001 | Thickness | Thickness of film | A | FLOAT | 4 Bytes
Between 10 and 399% | Tooling | the correction needed between the thickness of the layer deposited on the crystal monitor and the thickness of the film deposited on the substrate | % | INT | 4 Bytes
99.99 | Density | Density of material | g/cm^3 | FLOAT | 4 Bytes
0.10 to 10.00 | Z-ratio | Z-ratio of material | - | FLOAT | 4 Bytes
| Sampling frequency | Resolution of recorded rate and thickness data | 1s
| Time index | An integer index for the tooling, density and zratio entries |  | INT | 4 Bytes
General structure of the dataframes
3 Dataframes
| Dataframe1: Rate and thickness will be sent at each time step |  |  |  | Dataframe2: An entry is only populated when the value changes |  | Dataframe3: An entry is only populated when the value changes |  | Dataframe4: An entry is only populated when the value changes
| Time (s) | Rate (A/s) | Thickness (A) |  | Time INDEX | Tooling (%) | Time INDEX | Density (g/cm^3) | Time INDEX | Z-Ratio
I think the default sample rate is 1s but will confirm this | 0 | Rate1 | Thickness1 |  | t1 | Tooling1 | t1 | Density1 | t1 | Z-Ratio1
| 1 | Rate2 | Thickness2 |  | t2 | Tooling2 | t2 | Density2 | t2 | Z-Ratio2
| 2 | Rate3 | Thickness3 |  | . | . | . | . | . | .
| 3 | Rate4 | Thickness4 |  | . | . | . | . | . | .
| 4 | Rate5 | Thickness5 |  | . | . | . | . | . | .
| 5 | Rate6 | Thickness6 |  | tn | n | tn | n | tn | n
| 6 | Rate7 | Thickness7
| . | . | .
| . | . | .
| . | . | .
| n | n | n
What will actually be sent
| [[rate1,thickness1], [rate2,thickness2], …, [raten,thicknessn]] |  |  |  | [[Time_index1, tooling1], [Time_index2, tooling2], …, [Time_indexn, toolingn]] |  | [[Time_index1, density1], [Time_index2, density2], …, [Time_indexn, densityn]] |  | [[Time_index1, zratio1], [Time_index2, zratio2], …, [Time_indexn, zration]]
Commands from Host to SQM-160
| Command structure from host to SQM-160 | <Sync character><Length character><Message><CRC1><CRC2>
| Command structure from SQM-160 to host | <Sync character><Length character><Response Status character><Response Message><CRC1><CRC2>
| Command U | Read the status of the shutter
| Command structure: | U?
| Shutter status | 0 or 1. 0 closes the shutter. 1 opens the shutter
| Example status | !$U?(91)(84)
| Example received | !%A0(118)(135)
| Command L | Read the current rate for a sensor
| Command structure: | L<Sensor number>?
| Sensor Number | 1 to 6
| Example sent | !%L1?(133)(123)
| Example received | !*A_0.00_(91)(100)
| Command N | Read the current thickness for a sensor
| Command structure: | N<Sensor number>
| Sensor Number | 1 to 6
| Example sent | !$N1(93)(81)
| Example received | !+A_0.000_(74)(111)
| Command A | Read the density, tooling and z-ratio
| Example sent | !%A1?(46)(149)
| Example received | !0AFILM 1____1.00_100_1.000_33.380__0.000_0_1(79)(59)
| Structure of received ^^ | A<Film Number><Film name>_<Density>_<Film Tooling>_<Z-factor>_<Final Thickness>_<Thickness Setpoint>_<Time Setpoint>_<Sensor Average>
|  | DEC | HEX | BIN | CHAR
|  | 0 | 00000000 | 00000000 | #VALUE!
|  | 1 | 00000001 | 00000001 | 
|  | 2 | 00000002 | 00000010 | 
|  | 3 | 00000003 | 00000011 | 
|  | 4 | 00000004 | 00000100 | 
|  | 5 | 00000005 | 00000101 | 
|  | 6 | 00000006 | 00000110 | 
|  | 7 | 00000007 | 00000111 | 
|  | 8 | 00000008 | 00001000 | 
|  | 9 | 00000009 | 00001001
|  | 10 | 0000000A | 00001010
|  | 11 | 0000000B | 00001011
|  | 12 | 0000000C | 00001100
|  | 13 | 0000000D | 00001101
|  | 14 | 0000000E | 00001110 | 
|  | 15 | 0000000F | 00001111 | 
|  | 16 | 00000010 | 00010000 | 
|  | 17 | 00000011 | 00010001 | 
|  | 18 | 00000012 | 00010010 | 
|  | 19 | 00000013 | 00010011 | 
|  | 20 | 00000014 | 00010100 | 
|  | 21 | 00000015 | 00010101 | 
|  | 22 | 00000016 | 00010110 | 
|  | 23 | 00000017 | 00010111 | 
|  | 24 | 00000018 | 00011000 | 
|  | 25 | 00000019 | 00011001 | 
|  | 26 | 0000001A | 00011010 | 
|  | 27 | 0000001B | 00011011 | 
|  | 28 | 0000001C | 00011100
|  | 29 | 0000001D | 00011101
|  | 30 | 0000001E | 00011110
|  | 31 | 0000001F | 00011111
|  | 32 | 00000020 | 00100000
|  | 33 | 00000021 | 00100001 | !
|  | 34 | 00000022 | 00100010 | "
|  | 35 | 00000023 | 00100011 | #
|  | 36 | 00000024 | 00100100 | $
|  | 37 | 00000025 | 00100101 | %
|  | 38 | 00000026 | 00100110 | &
|  | 39 | 00000027 | 00100111 | '
|  | 40 | 00000028 | 00101000 | (
|  | 41 | 00000029 | 00101001 | )
|  | 42 | 0000002A | 00101010 | *
|  | 43 | 0000002B | 00101011 | +
|  | 44 | 0000002C | 00101100 | ,
|  | 45 | 0000002D | 00101101 | -
|  | 46 | 0000002E | 00101110 | .
|  | 47 | 0000002F | 00101111 | /
|  | 48 | 00000030 | 00110000 | 0
|  | 49 | 00000031 | 00110001 | 1
|  | 50 | 00000032 | 00110010 | 2
|  | 51 | 00000033 | 00110011 | 3
|  | 52 | 00000034 | 00110100 | 4
|  | 53 | 00000035 | 00110101 | 5
|  | 54 | 00000036 | 00110110 | 6
|  | 55 | 00000037 | 00110111 | 7
|  | 56 | 00000038 | 00111000 | 8
|  | 57 | 00000039 | 00111001 | 9
|  | 58 | 0000003A | 00111010 | :
|  | 59 | 0000003B | 00111011 | ;
|  | 60 | 0000003C | 00111100 | <
|  | 61 | 0000003D | 00111101 | =
|  | 62 | 0000003E | 00111110 | >
|  | 63 | 0000003F | 00111111 | ?
|  | 64 | 00000040 | 01000000 | @
|  | 65 | 00000041 | 01000001 | A
|  | 66 | 00000042 | 01000010 | B
|  | 67 | 00000043 | 01000011 | C
|  | 68 | 00000044 | 01000100 | D
|  | 69 | 00000045 | 01000101 | E
|  | 70 | 00000046 | 01000110 | F
|  | 71 | 00000047 | 01000111 | G
|  | 72 | 00000048 | 01001000 | H
|  | 73 | 00000049 | 01001001 | I
|  | 74 | 0000004A | 01001010 | J
|  | 75 | 0000004B | 01001011 | K
|  | 76 | 0000004C | 01001100 | L
|  | 77 | 0000004D | 01001101 | M
|  | 78 | 0000004E | 01001110 | N
|  | 79 | 0000004F | 01001111 | O
|  | 80 | 00000050 | 01010000 | P
|  | 81 | 00000051 | 01010001 | Q
|  | 82 | 00000052 | 01010010 | R
|  | 83 | 00000053 | 01010011 | S
|  | 84 | 00000054 | 01010100 | T
|  | 85 | 00000055 | 01010101 | U
|  | 86 | 00000056 | 01010110 | V
|  | 87 | 00000057 | 01010111 | W
|  | 88 | 00000058 | 01011000 | X
|  | 89 | 00000059 | 01011001 | Y
|  | 90 | 0000005A | 01011010 | Z
|  | 91 | 0000005B | 01011011 | [
|  | 92 | 0000005C | 01011100 | \
|  | 93 | 0000005D | 01011101 | ]
|  | 94 | 0000005E | 01011110 | ^
|  | 95 | 0000005F | 01011111 | _
|  | 96 | 00000060 | 01100000 | `
|  | 97 | 00000061 | 01100001 | a
|  | 98 | 00000062 | 01100010 | b
|  | 99 | 00000063 | 01100011 | c
|  | 100 | 00000064 | 01100100 | d
|  | 101 | 00000065 | 01100101 | e
|  | 102 | 00000066 | 01100110 | f
|  | 103 | 00000067 | 01100111 | g
|  | 104 | 00000068 | 01101000 | h
|  | 105 | 00000069 | 01101001 | i
|  | 106 | 0000006A | 01101010 | j
|  | 107 | 0000006B | 01101011 | k
|  | 108 | 0000006C | 01101100 | l
|  | 109 | 0000006D | 01101101 | m
|  | 110 | 0000006E | 01101110 | n
|  | 111 | 0000006F | 01101111 | o
|  | 112 | 00000070 | 01110000 | p
|  | 113 | 00000071 | 01110001 | q
|  | 114 | 00000072 | 01110010 | r
|  | 115 | 00000073 | 01110011 | s
|  | 116 | 00000074 | 01110100 | t
|  | 117 | 00000075 | 01110101 | u
|  | 118 | 00000076 | 01110110 | v
|  | 119 | 00000077 | 01110111 | w
|  | 120 | 00000078 | 01111000 | x
|  | 121 | 00000079 | 01111001 | y
|  | 122 | 0000007A | 01111010 | z
|  | 123 | 0000007B | 01111011 | {
|  | 124 | 0000007C | 01111100 | ¦